Background

Colorimeters and spectrophotometers are important pieces of laboratory equipment and are used in many experiments in schools and colleges.  A quick look through the Course Specifications for the Revised and CfE Highers and Advanced Highers flags up a number of opportunities for the use of colorimeters and spectrophotometers.  It is not that common for the reliability of colorimeters to be tested and we offer this short practical as a way of testing their accuracy, reliability and precision.

Method

| Set the filter wheel to 440 nm Switch on the colorimeter Place an empty cuvette into the sample holder and press the ‘R’ button to zero the colorimeter. Note the direction of the light beam - ensure cuvette is in the correct orientation. |  | Direction of beam |
| --- | --- | --- |

- Press the ‘T’ button and record the absorbance.

- Cut 5 pieces of neutral density filter each with dimensions 3 cm x 0.8 cm.

- Place (tweezers may help) a piece of neutral density filter into the cuvette and make sure it is vertical.

- Press the ‘T’ button and record the absorbance.  Add another piece of neutral density filter and measure the absorbance.  Repeat until you have a total of 6 data points.

- Repeat steps 3-7 with the filter wheel at 520 nm.

- Plot your results

Equipment Needed

- Colorimeter (in the experiment today we will use a WPA CO7500 colorimeter but in principle this experiment could be repeated with any colorimeter).

- Scissors, ruler, tweezers.

- Cuvette (clean, empty).

- Sheet of neutral density filter.  This filter is manufactured such that a single thickness allows 50% of light to be transmitted.

- Graph paper.

# Text boxes

How good is the colorimeter?